Supplementary material

for

LC-MS/MS quantification of vitamin K1 after simple precipitation of protein from low volume of serum

Hai-Bo Wang†a, Rui Peng†a, Juan Le†a and Shao-Ting Wang#a

† These authors contributed equally to the present work.

# To whom correspondence should be addressed:

a Department of Clinical Laboratory, Renmin Hospital of Wuhan University, Wuhan 430060, China

Fax: +86-027-88041911; E-mail: shaotingw@whu.edu.cn

Supplemental Table S1. The flow program for LC separation.

| Time (min) | Flow rate (mL/min) | Flow direction |
| --- | --- | --- |
| 0-1.7 | 0.5 | To the waste |
| 1.7-2.7 | 0.5 | To the ion source |
| 2.7-2.8 | 0.5 → 1.0 | To the waste |
| 2.8-5.8 | 1.0 | To the waste |
| 5.8-6.0 | 1.0 → 0.5 | To the waste |

Supplemental Table S2. The MS parameters of K1 and K1-d7.

| Targets | Precursor ion (m/z) | Product ion (m/z) | CE (V) |
| --- | --- | --- | --- |
| K1 a | 451.3 | 187.2 | -26 |
| K1 b | 451.3 | 171.2 | -37 |
| K1-d7 | 458.3 | 194.2 | -26 |

a Quantifier ion.

b Qualifier ion.

Supplemental Table S3. The releasing-efficiency of K1 after protein precipitation by methanol, ethanol, isopropanol or acetonitrile (as described in EXPERIMENT I and Equation 1).

| Volume ratio of serum to precipitator | Releasing-efficiency of K1 (%) ± SD after PP |  |  |  |
| --- | --- | --- | --- | --- |
|  | Methanol | Ethanol | Isopropanol | Acetonitrile |
| 1:2 | 4.0 ± 0.4 | 27.7 ± 2.2 | 68.1 ± 6.8 | 9.5 ± 0.9 |
| 1:3 | 7.0 ± 0.8 | 65.6 ± 3.3 | 92.3 ± 2.8 | 10.0 ± 1.7 |
| 1:4 | 24.3 ± 1.9 | 74.0 ± 2.2 | 94.4 ± 6.2 | 6.7 ± 0.8 |
| 1:8 | 64.2 ± 4.5 | 82.2 ± 5.8 | 90.5 ± 3.6 | 5.1 ± 1.1 |
| 1:16 | 84.1 ± 5.9 | 89.8 ± 5.4 | 91.3 ± 7.3 | 16.4 ± 4.3 |

Supplemental Table S4. The releasing-efficiency of K1 regaining from washing the precipitate with ethanol after protein precipitation by methanol, ethanol, isopropanol or acetonitrile (as described in EXPERIMENT I and Equation 2).

| Volume ratio of serum to precipitator | Releasing-efficiency of K1 (%) ± SD after washing the precipitate |  |  |  |
| --- | --- | --- | --- | --- |
|  | Methanol | Ethanol | Isopropanol | Acetonitrile |
| 1:2 | 68.8 ± 11.0 | 38.0 ± 2.3 | 9.2 ± 0.6 | 7.4 ± 1.3 |
| 1:3 | 73.3 ± 5.9 | 17.8 ± 1.6 | 5.3 ± 0.4 | 5.0 ± 0.8 |
| 1:4 | 52.4 ± 3.1 | 10.7 ± 1.3 | 7.8 ± 0.2 | 6.1 ± 1.8 |
| 1:8 | 14.9 ± 1.2 | 6.4 ± 0.8 | 7.7 ± 0.5 | 8.9 ± 5.2 |
| 1:16 | 4.9 ± 0.3 | 4.3 ± 0.1 | 6.6 ± 0.3 | 4.3 ± 1.0 |

Supplemental Table S5. The releasing-efficiency of K1 after protein precipitation by methanol, ethanol, isopropanol or acetonitrile coupled with liquid-liquid extraction with n-hexane (as described in EXPERIMENT II and Equation 3).

| Volume ratio of serum to precipitator | Releasing-efficiency of K1 (%) ± SD after PP&LLE |  |  |  |
| --- | --- | --- | --- | --- |
|  | Methanol | Ethanol | Isopropanol * | Acetonitrile |
| 1:2 | 20.8 ± 5.6 | 57.4 ± 6.4 | - | 10.0 ± 2.4 |
| 1:3 | 50.7 ± 8.3 | 70.5 ± 8.9 | - | 7.4 ± 2.0 |
| 1:4 | 72.7 ± 11.0 | 73.3 ± 8.7 | - | 15.0 ± 3.9 |
| 1:8 | 74.1 ± 7.4 | 65.2 ± 10.8 | - | 12.7 ± 1.0 |
| 1:16 | 70.2 ± 14.2 | 38.6 ± 5.2 | - | 11.3 ± 2.2 |

* As isopropanol is miscible with n-hexane, the obtained upper phase after LLE needs too much time to dry under general circumstance. In result, this group of samples was not subjected to further LC-MS/MS analysis.

Supplemental Table S6. The releasing-efficiency of K1 regaining from washing the precipitate with ethanol after protein precipitation by methanol, ethanol, isopropanol or acetonitrile coupled with liquid-liquid extraction with n-hexane (as described in EXPERIMENT II and Equation 2).

| Volume ratio of serum to precipitator | Releasing-efficiency of K1 (%) ± SD after washing the precipitate after PP&LLE |  |  |  |
| --- | --- | --- | --- | --- |
|  | Methanol | Ethanol | Isopropanol | Acetonitrile |
| 1:2 | 47.5 ± 6.0 | 18.8 ± 0.9 | 3.8 ± 1.3 | 5.5 ± 1.2 |
| 1:3 | 22.5 ± 2.1 | 6.3 ± 1.3 | 1.3 ± 1.0 | 5.0 ± 1.3 |
| 1:4 | 10.0 ± 1.4 | 1.3 ± 0.4 | 2.5 ± 0.5 | 5.5 ± 1.8 |
| 1:8 | 3.8 ± 0.2 | 2.5 ± 0.6 | 2.6 ± 0.3 | 2.5 ± 1.1 |
| 1:16 | 2.5 ± 0.5 | 1.3 ± 0.2 | 1.3 ± 0.3 | 4.3 ± 0.9 |

Supplemental Table S7. The impact of injected serum per run on ion suppression effect for K1 and IS (relative to GROUP 1) after protein precipitation coupled with liquid-liquid extraction pretreatment (as described in EXPERIMENT III and Equation 5 & 6).

| GROUP X | Injected serum per run (μL) | Ion suppression effect on K1 (%) | Ion suppression effect on IS (%) |
| --- | --- | --- | --- |
| GROUP 1 | 1 | 0 | 0 |
| GROUP 2 | 2 | 4.4 | 6.7 |
| GROUP 3 | 5 | 31.7 | 25.0 |
| GROUP 4 | 10 | 51.1 | 45.7 |
| GROUP 5 | 20 | 74.3 | 86.7 |
| GROUP 6 | 50 | 92.2 | 95.4 |

Supplemental Table S8. The impact of injected serum per run on detected peak area of K1 after protein precipitation coupled with liquid-liquid extraction pretreatment (as described in EXPERIMENT III).

| GROUP X | Injected serum per run (μL) | Detected peak area of K1 | Theoretical peak area of K1 calculated from GROUPT 1 |
| --- | --- | --- | --- |
| GROUP 1 | 1 | 8900 | 8900 |
| GROUP 2 | 2 | 17012 | 17800 |
| GROUP 3 | 5 | 30373 | 44500 |
| GROUP 4 | 10 | 43526 | 89000 |
| GROUP 5 | 20 | 45686 | 178000 |
| GROUP 6 | 50 | 34565 | 445000 |

Supplemental Table S9. The linear regressions built from three different blank matrixes.

| Blank matrix | Linear regression |  |  |
| --- | --- | --- | --- |
|  | Slope | Intercept | R2 |
| I | 0.996 | 0.024 | 0.996 |
| II | 1.012 | -0.013 | 0.998 |
| III | 1.030 | -0.010 | 0.999 |

Supplemental Table S10. The method accuracy (as apparent recovery) and imprecision (as CV) for K1 determination at four concentration levels.

|  | Apparent recovery of K1, % (CV, %) |  |  |  |
| --- | --- | --- | --- | --- |
|  | High-level | Medium-level | Low-level | LLOQ-level |
| Intra-day (n=6) | 99.4 (4.2) | 102.8 (5.4) | 91.5 (8.6) | 104.1 (10.6) |
| Inter-day (n=10) | 97.2 (6.7) | 94.2 (7.7) | 92.3 (8.1) | 112.2 (13.5) |

Supplemental Table S11. The concentration of K1 detected in one hundred random samples.

| Sample ID | K1 concentration (ng/mL) | Sample ID | K1 concentration (ng/mL) | Sample ID | K1 concentration (ng/mL) |
| --- | --- | --- | --- | --- | --- |
| 1 | 0.09 | 35 | 0.72 | 69 | 0.56 |
| 2 | 0.12 | 36 | 0.12 | 70 | 0.43 |
| 3 | 0.15 | 37 | 0.30 | 71 | 0.11 |
| 4 | 0.30 | 38 | 0.17 | 72 | 0.10 |
| 5 | 0.52 | 39 | 2.40 | 73 | 0.23 |
| 6 | 0.37 | 40 | 0.89 | 74 | 0.99 |
| 7 | 0.29 | 41 | 0.76 | 75 | 2.94 |
| 8 | 0.42 | 42 | 1.37 | 76 | 0.57 |
| 9 | 0.41 | 43 | 0.74 | 77 | 0.35 |
| 10 | 0.35 | 44 | 0.42 | 78 | 0.46 |
| 11 | 0.17 | 45 | 0.28 | 79 | 0.40 |
| 12 | 0.37 | 46 | 0.27 | 80 | 0.36 |
| 13 | 0.12 | 47 | 0.47 | 81 | 0.41 |
| 14 | 0.29 | 48 | 0.32 | 82 | 1.23 |
| 15 | 0.30 | 49 | 0.65 | 83 | 0.91 |
| 16 | 0.11 | 50 | 0.63 | 84 | 0.51 |
| 17 | 0.28 | 51 | 0.30 | 85 | 1.40 |
| 18 | 0.41 | 52 | 0.38 | 86 | 0.75 |
| 19 | 0.60 | 53 | 0.67 | 87 | 0.32 |
| 20 | 0.63 | 54 | 0.31 | 88 | 1.47 |
| 21 | 0.14 | 55 | 0.12 | 89 | 0.28 |
| 22 | 0.07 | 56 | 0.71 | 90 | 0.73 |
| 23 | 0.25 | 57 | 0.09 | 91 | 0.89 |
| 24 | 0.17 | 58 | 0.27 | 92 | 1.17 |
| 25 | 0.54 | 59 | 0.61 | 93 | 0.36 |
| 26 | 0.73 | 60 | 0.18 | 94 | 0.84 |
| 27 | 0.31 | 61 | 0.29 | 95 | 1.28 |
| 28 | 1.43 | 62 | 0.32 | 96 | 1.82 |
| 29 | 0.27 | 63 | 0.40 | 97 | 0.56 |
| 30 | 1.67 | 64 | 0.16 | 98 | 1.23 |
| 31 | 0.78 | 65 | 0.34 | 99 | 0.90 |
| 32 | 0.72 | 66 | 0.18 | 100 | 1.40 |
| 33 | 1.08 | 67 | 0.34 |  |  |
| 34 | 0.34 | 68 | 0.18 |  |  |

Supplemental Figure S1. The residual K1 in C0 from (A) BLANK MATRIX I and (B) BLANK MATRIX II.